## Odabiri (hidden)
Odgovorne osobe za pokretanje jednostavne nabave (2020) | Organizacijska jedinica | Rok, način i uvjeti plaćanja: | Mjesto traženja zahtjeva |  | Odabir Izuzeća | Vrsta Ugovora o jednostavnoj nabavi |  |  |  | Točka Pravilnika | Izvor financiranja | Prodekan za poslovanje | Odobrenje zahtjeva (2020)
Meaški, Hrvoje | Zavod za geotehniku | Gotovina (refundacija) | Varaždinu | ROBA | DA | Jednokratna nabava | Najniža cijena | 30 dana | Postupak jednostavne nabave roba. | 1 (npr. zbog zaštite zdravlja i sigurnosti zaposlenika, nastanka znatne štete na materijalnoj imovini te značajan utjecaj na obavljanje poslova iz djelokruga Naručitelja) | Troškovi poslovanja - proračunski | Biondić, Ranko | Meaški, Hrvoje | roba
Biondić, Ranko | Zavod za hidrotehniku | Po ponudi (predujam) | Zagrebu | USLUGA | NE | Godišnji ugovor o jednostavnoj nabavi | Ekonomski najpovoljnija ponuda | 60 dana | Postupak jednostavne nabave usluga. | 2 (npr. autorska i slična djela, prijava patenata, objava znanstvenih članaka, kotizacije za konferencije, skupove i seminare, pretplate na časopise, akreditacije laboratorija i slično ) | Troškovi poslovanja - fakultetska sredstva |  |  | usluga
Grčić, Ivana | Zavod za inženjerstvo okoliša | Transakcijski račun |  | RADOVA |  | Okvirni sporazum o jednostavnoj nabavi do dvije godine |  | 90 dana | Postupak jednostavne nabave radova. | 3 (npr. odvjetničke i javnobilježničke usluge, konzultantske usluge, usluge obrazovanja i stručnog osposobljavanja, zdravstvene usluge, usluge javnog prijevoza, usluge kupnje avionskih karata, usluge hotelskog smještaja, reprezentacije i slično) | Stručni projekt |  |  | radova
Kovač, Sanja | Zavod za opće znanosti | Kompenzacija | [5 empty] | 120 dana |  | 4 (kod nabava koje provodi Središnji državni ured za središnju javnu nabavu) | Znanstveni projekt
Leskovar, Katica | Računovodstvo | [8 empty] | 5 (kada u postupku prikupljanja ponuda nije dostavljena niti jedna valjana ponuda, a postupak jednostavne nabave se ponavlja) | Sredstva na osobnoj kartici
Anić Vučinić, Aleksandra | Tajništvo | [8 empty] | 6 (kod nabava koje su propisane, kao izuzeća, člancima 29.-35. ZJN 2016) | Zavod za geotehniku
Kapelj, Sanja | [10 empty] | Zavod za geotehniku - GEOLAB
Strelec, Stjepan | [10 empty] | Zavod za hidrotehniku
Stančić, Zvjezdana | [10 empty] | Zavod za hidrotehniku - LGO
Flajšman, Mirjana | [10 empty] | Zavod za inženjerstvo okoliša
Zavod za inženjerstvo okoliša - LIO
Zavod za opće znanosti
Refundacija organizatora / vanjske ustanove

## Prilog 1-Troskovnik-Cost Sheet
Prilog 1. Troškovnik - podvodni fluorimetar / Annex 1: Cost Sheet - Submersible fluorimeter
Sveučilište u Zagrebu, Geotehnički fakultet / University of Zagreb, Faculty of Geotechnical Engineering
Hallerova aleja 7, 42000 Varaždin / Hallerova aleja 7, 42000 Varaždin
EVIDENCIJSKI BROJ NABAVE: Jednostavna nabava 161/2021 / PROCUREMENT RECORD NUMBER: Simple procurement 161/2021
NABAVA OPREME: Podvodni fluorimetar / PURCHASE OF EQUIPMENT: Submersible fluorimeter
R. BR. | STAVKA/SPECIFIKACIJA | Jedinica | Količina | Jedinična cijena bez PDV-a | Ukupna cijena bez PDV-a | Zadovoljava zahtjev minimalne tehničke specifikacije, DA / NE
No. | ITEM / SPECIFICATION | Unit | Quantity | Unit price without tax | Total price without tax | Meets the requirement of a minimum technical specification, YES / NO
1 | C3 Submersible Fluorometer Package or better, configured as follows (1.1.-1.8): | kom / piece | 1
1.1 | Instrument package should includes interface cable for PC, 12V power supply, factory-installed temperature sensor, communication and integration cable ("Y" style USB to male & female DB-9), antifouling copper tape, one 8-pin female end plug and software, European Line Cord for power supply
1.2 | Fluorescein Optics with following specifications: MDL 0.01 ppb; linear range: 0-500 ppb; LED (CWL): 460 nm; Excitation: 400/150 nm; Emission: 545/28 nm
1.3 | Rhodamine WT Optics with following specifications: MDL 0.01 ppb; linear range: 0-1000 ppb; LED (CWL): 530 nm; Excitation: 535/60 nm; Emission: 590-715 nm
1.4 | Optical Brighteners Optic with following specifications: MDL 0.6 ppb; linear range: 0-2500 ppb; LED (CWL): 365 nm; Excitation: 325/120 nm; Emission: 445/15 nm
1.5 | Pressure Sensor (should be factory Installed)
1.6 | C-SOFT Firmware (for intuitive calibration, data logging, set up, and download functions of C3 Submersible Fluorometer)
1.7 | Minimum 1-Year Warranty
2 | Rechargable Battery Pack to provide self-contained use of C3 Submersible Fluorometer (should contains a rechargeable lithium-ion battery and battery pack charger) Minimum capacity: 2,200 mAH or higher | kom / piece | 2
3 | C3 Battery Bracket (required for use with Standard Battery Pack, Analog Output Adapter, and Digital Converter if C3 is not in C-ray Towed Deployment Body) | kom / piece | 1
4 | C3 Shade Cap (prevent damage from of the optical head and minimizes effects from ambient light) | kom / piece | 1
5 | Calibration services | usluga / service | 1
6 | Transport and Delivery costs | usluga / service | 1
UKUPNA CIJENA BEZ PDV-a / TOTAL PRICE WITHOUT TAX
PDV 25% / Value-added tax (VAT) 25%
UKUPNA CIJENA S PDV-om / TOTAL PRICE WITH VAT